## Slide 1: SDVM Mask
FINAL Report
Presentation
1
Social Distancing and Vital Monitoring Mask

## Slide 2: Contents
Background
Problem Statement
Objectives of Project
Implementation of Design
  Block Diagram/Flowchart
  Testing
  Final Product
  Application
Results and Discussion
Impact of Solutions
Conclusions
References
2

## Slide 3: Current situation
TWO meter
SOCIAL DISTANCING
Wearing a mask
Customers
Who?
Citizens of COVID-19 affected countries
What?
Help them live out their day-to-day lives.
3
1. Background
GETTING VACCINE
Notes: 210 countries around the world have reported more than three million infected people including more than two hundred thousand deaths [1]
Current Situation: Washing hands

## Slide 4: 2. Problem statement
4
The spreading of the contagious COVID-19 disease worldwide is alarmingly fast, and there are instances of infected individuals showing delayed signs of symptoms.
There are difficulties containing the virus and measures suggested such as social distancing to mitigate the spread
Virus can affect the human body and cause a range of symptoms, in which initial information and evaluation of vital signs regarding a possible infection are crucial

## Slide 5: 3. Objectives of Project
01
03
02
04
Design a social distancing and vital signs monitoring lightweight mask.
To measure the user's vital signs (i.e., temperature and oxygen saturation)
To provide warnings to alert the user when another person is too close to maintain proper social distancing
To build an application that tracks the users’ vital signs and alarms them of any abnormalities
5
05
Model the mask structure with a cleanable front material and detachable straps that hold the sensors

## Slide 6: 04.
IMPLEMENTATION

## Slide 7: FLOW Chart

## Slide 8: Testing Components

## Slide 9
1. Testing GY-906 Sensor

## Slide 10
2 . Testing MAX30100 Sensor

## Slide 11
3. Testing PIR Sensor

## Slide 12
4. Bluetooth Testing

## Slide 13: FINAL PRODUCT
PIR Sensor at the front of mask
Temperature Sensor and Pulse Oximeter on front of mask from inside
Notes: Explain how its worn
Mention that both temperature and pulse oximeter are able to read through cloth covering
PIR: Specific maximum range, which is 2 meters , and angle (<120)

## Slide 14: FINAL PRODUCT
Battery, Arduino’s, and Bluetooth Module on back piece.
PIR Sensor at the front of mask
Notes: Mention reason for two Arduinos in final product.
Whole back piece with straps is detachable so that front material can get cleaned.

## Slide 15: APPLICATION
MIT App Inventor used to build application
Application displays notifications in case:
  Temperature is high
  Oxygen saturation is low
  Social distancing is compromised

## Slide 16: Results and discussion 05.
Testing was conducted on 5 different volunteers!

## Slide 17: Temperature Sensor Readings
Test Subject | Age | MLX90614 Body Temperature (C) | Thermometer (C) | Difference
Person 1 | 22 | 36.55 | 37.05 | 0.5
Person 2 | 45 | 34.91 | 35.04 | 0.13
Person 3 | 24 | 35.33 | 35.45 | 0.12
Person 4 | 18 | 35.67 | 35.74 | 0.07
Person 5 | 12 | 35.73 | 36 | 0.27
Table – Body temperature measurement results using the MLX90614 sensor

## Slide 18: Temperature Sensor Readings
Reading increase highly and takes from 500 to 1000 milliseconds to be nearly stable!
Figure - Plot of temperature readings for each test subject
Person 5
Person 4
Person 3
Person 2
Person 1

## Slide 19: Blood Oxygen Saturation Sensor Readings
Table - Oxygen saturation measurement results using the MAX30100 sensor
Test Subject | Age | MAX30100 Oxygen Saturation (%) | Samsung Oxygen Saturation (%) | Difference
Person 1 | 22 | 98 | 98 | 0
Person 2 | 45 | 95 | 99 | 4
Person 3 | 24 | 98 | 98 | 1
Person 4 | 18 | 96 | 98 | 2
Person 5 | 12 | 98 | 99 | 1

## Slide 20: Social Distancing Sensor Readings
The sensor is set to stay at HIGH for 20 seconds as long as the person is still standing in front of the user!
Figure - Plot showing the relation between time and distance for the PIR sensor

## Slide 21: EVALUATION OF SOLUTIONS
04
05
01
02
03
The optimal power usage was satisfied
The mask is environmentally friendly
The SDVM mask is a lightweight disposable mask with a cleanable front material structure
The mask can track the user's social distancing and vital signs.
The mask is provided with an application that tracks the users’ vital signs and warns them of any abnormalities

## Slide 22
Washable
Recyclable
Reliable power source.
- Affect the country’s economy positively.
Decrease the load on the health sector.
Make the routine checkups for healthcare workers more constant and time-efficient.
Impacts of solutions
Financial
Safety
Social
Environmental

## Slide 23: Conclusions
The prototype was able to model the desired outputs of the system and meet customer needs and requirements. The main objectives of the prototype were for it to be able to help the user monitor their vital signs and maintain proper social distancing. The application also serves its purpose in being a user friendly and simple manner.

## Slide 24
THANKS!
Do you have any questions?
18

## Slide 25: References
19
 | [1] C. T. Nguyen, "Enabling and Emerging Technologies for Social Distancing: A Comprehensive Survey and Open Problems," p. 43, 2020
 | [2] T. TAMURA, M. HUANG and T. TOGAWA, "Thermometers, Current Developments in Wearable," Advanced Biomedical Engineering, .p. 13, 2018 [3] Y. Khan, A. E. Ostfeld, C. M. Lochner, A. Pierre and A. C. Arias, "Monitoring of Vital Signs with Flexible and Wearable," Advanced Material, 2016. [4] S. Bian, B. Zhou and P. Lukowicz, "Social Distance Monitor with a Wearable Magnetic Field Proximity Sensor," MDPI Open Access Journals, vol. 20, no. 18, p. 5101, September 2020. [5] Z. Wang, Z. Yang and T. Dong, "A Review of Wearable Technologies for Elderly Care that Can Accurately Track Indoor Position," MDPI Open Access Journals, vol. 17, p. 341, 2017. [6] R. Effing, R. J.Brouwer, A. Iliev, T. Teunissen and F. Wijnhoven, "WHERE TO GO? SMART GUIDANCE BASED ON IOT SENSOR-DATA," International Conference ICT, Soceity, and Human Beings, 2020. [7] A. Alarifi, "UltraWideband Indoor Positioning Technologies: Analyis and Recent Advances," MDPI Open Access Journal, vol. 16, p. 707, 2016. [8] C. T. Nguyen, "Enabling and Emerging Technologies for Social Distancing: A Comprehensive Survey and Open Problems," p. 43, 2020. [9] E. Chapin, "University of Kentucky, College of Engineering," 14 May 2020. [Online]. Available: https://www.engr.uky.edu/research-faculty/research/2020/developingmask-captures-and-deactivates-covid-19-dibakar. [Accessed 5 October 2020].

## Slide 26: References
20
 | [10] S. Kelly, "Cornell Chronicle," 6 May 2020. [Online]. Available: https://news.cornell.edu/stories/2020/05/student-team-designs-smart-mask-monitorsvital-signs. [Accessed 5 October 2020]. [11] A. Bendix, "Business Insider," 13 May 2020. [Online]. Available: https://www.businessinsider.com/coronavirus-face-mask-light-up-screening-tooltest-2020-5. [Accessed 5 October 2020]. [12] "Carebook," [Online]. Available: https://carebook.com/my-vitals-covid-19/. [Accessed 13 October 2020]. [13] C. Criddle, "BBC Tech," 9 September 2020. [Online]. Available: https://www.bbc.com/news/technology-54091539. [Accessed 6 October 2020].